ISO 14971:2019 Risk Management Training

Post-Training Knowledge Quiz

Instructions: For each question, select the single best answer from the five options provided.

Name:

Date:

Question 1

A reusable surgical instrument is inadequately sterilized between procedures, leaving residual biological contamination on the device surface. In ISO 14971 terminology, which of the following correctly identifies the hazardous situation in this scenario?

A)  Microbiological hazard

B)  Surgical site infection

C)  Inadequate reprocessing procedure

D)  Patient’s surgical wound exposed to contaminated instrument surfaces

E)  Non-compliance with reprocessing instructions

Question 2

The manufacturer shall identify risk control measures using the following priority order:

A)  1. Protective measures > 2. Inherent safety by design > 3. Information for safety

B)  1. Information for safety > 2. Protective measures > 3. Inherent safety by design

C)  1. Inherent safety by design > 2. Protective measures > 3. Information for safety

D)  1. Inherent safety by design > 2. Information for safety > 3. Protective measures

E)  1. Protective measures > 2. Information for safety > 3. Inherent safety by design

Question 3

After implementing all practical risk control measures for an identified risk, the residual risk is judged not acceptable. According to ISO 14971, what is the next required step?

A)  Remove the feature causing the hazard from the device

B)  Accept the risk and document the justification

C)  Conduct a benefit-risk analysis to determine if the benefits outweigh the residual risk

D)  Add a warning to the Instructions for Use and accept the risk

E)  Escalate the decision to the notified body for approval

Question 4

ISO 14971 requires risk management activities to be conducted throughout which of the following?

A)  Only during the design and development phase

B)  From initial concept through to post-market surveillance, covering the entire device lifecycle

C)  Only until the device receives regulatory approval or CE marking

D)  During design, manufacturing, and the first year of post-market experience

E)  Only when triggered by a complaint, incident, or audit finding

Question 5

A manufacturer identifies a new hazard through analysis of post-market surveillance data. According to ISO 14971, what must the manufacturer do?

A)  Document the hazard in the next periodic safety update report only

B)  Add the new hazard to the risk management file and re-evaluate risks, updating risk controls if necessary

C)  Notify the competent authority immediately before taking any other action

D)  Defer the analysis until the next scheduled risk management review

E)  Conduct a new clinical evaluation before updating the risk management file

Correct answers

Question 1

A reusable surgical instrument is inadequately reprocessed between procedures, leaving residual biological contamination on the device surface. In ISO 14971 terminology, which of the following correctly identifies the hazardous situation in this scenario?

A)  Microbiological hazard

B)  Surgical site infection

C)  Inadequate reprocessing procedure

D)  Patient’s surgical wound exposed to contaminated instrument surfaces

E)  Non-compliance with reprocessing instructions

Correct answer: D

Rationale: A hazardous situation requires a person to be exposed to a hazard. Option A describes the hazard itself, B describes the harm, C and E describe contributing factors or root causes. Only D identifies the circumstance where a patient is exposed to a hazard.

Question 2

The manufacturer shall identify risk control measures using the following priority order:

A)  1. Protective measures > 2. Inherent safety by design > 3. Information for safety

B)  1. Information for safety > 2. Protective measures > 3. Inherent safety by design

C)  1. Inherent safety by design > 2. Protective measures > 3. Information for safety

D)  1. Inherent safety by design > 2. Information for safety > 3. Protective measures

E)  1. Protective measures > 2. Information for safety > 3. Inherent safety by design

Correct answer: C

Rationale: ISO 14971 Clause 7.1 specifies the priority order: (1) inherent safety by design, (2) protective measures in the medical device itself or in the manufacturing process, and (3) information for safety. This reflects the principle that eliminating a hazard is preferable to protecting against it, which is preferable to warning about it.

Question 3

After implementing all practical risk control measures for an identified risk, the residual risk is judged not acceptable. According to ISO 14971, what is the next required step?

A)  Remove the feature causing the hazard from the device

B)  Accept the risk and document the justification

C)  Conduct a benefit-risk analysis to determine if the benefits outweigh the residual risk

D)  Add a warning to the Instructions for Use and accept the risk

E)  Escalate the decision to the notified body for approval

Correct answer: C

Rationale: ISO 14971 Clause 7.4 requires that when residual risk exceeds acceptability criteria after all practical risk controls have been applied, the manufacturer must conduct a benefit-risk analysis. Only if the benefits outweigh the residual risk can the risk be accepted. Simply documenting a justification (B) or adding a warning (D) without this analysis is insufficient.

Question 4

ISO 14971 requires risk management activities to be conducted throughout which of the following?

A)  Only during the design and development phase

B)  From initial concept through to post-market surveillance, covering the entire device lifecycle

C)  Only until the device receives regulatory approval or CE marking

D)  During design, manufacturing, and the first year of post-market experience

E)  Only when triggered by a complaint, incident, or audit finding

Correct answer: B

Rationale: ISO 14971 Clause 4.1 establishes that the risk management process applies across the entire lifecycle of the medical device - from initial concept, through design, manufacturing, and post-market phases. Risk management is a continuous, proactive process, not a phase-limited or reactive activity.

Question 5

A manufacturer identifies a new hazard through analysis of post-market surveillance data. According to ISO 14971, what must the manufacturer do?

A)  Document the hazard in the next periodic safety update report only

B)  Add the new hazard to the risk management file and re-evaluate risks, updating risk controls if necessary

C)  Notify the competent authority immediately before taking any other action

D)  Defer the analysis until the next scheduled risk management review

E)  Conduct a new clinical evaluation before updating the risk management file

Correct answer: B

Rationale: ISO 14971 Clause 10.3-10.4 requires that when new hazards or hazardous situations are identified from production and post-production information, the manufacturer must update the risk management file, assess the new risks, and implement additional risk controls as needed. This is an ongoing obligation, not something deferred to periodic reviews.